INSOUL BIOMECHANICS & SENSOR SYSTEMS

Ambulatory Gait Analysis & Biomechanical Algorithms

Engineering Reference & Mathematical Foundations: Ground Reaction Force (vGRF), Temporal Sub-Phasing, Bilateral Gait Symmetry, and Inertial Kinematics

| PLATFORM: ESP32 + Flutter | SAMPLING RATE: 50 Hz Real-Time Telemetry | SENSORS: Bilateral FSRs + 6-Axis IMU |
| --- | --- | --- |

1. Apparatus & Sensor Architecture

The InSoul system operates as an ambulatory, bilateral smart insole telemetry platform communicating over Bluetooth Low Energy (BLE). The sensory architecture captures vertical ground reaction forces and segment kinematic dynamics continuously.

• Heel Force-Sensitive Resistor (F_heel): Measures shock absorption, initial heel strike contact, and early stance loading response.

• Forefoot / Toe FSR (F_toe): Measures forefoot weight transfer, terminal stance stability, and propulsive pre-swing push-off.

• Telemetry Sample Rate: 50 Hz. High-frequency ADC sampling transmitting discrete real-time frames every 20 milliseconds.

• 6-Axis Inertial Measurement Unit (MPU-6050): 3-axis Accelerometer (±2g dynamic range, sensitivity 16,384 LSB/g) and 3-axis Gyroscope (±250°/s dynamic range, sensitivity 131 LSB/(°/s)).

• Embedded Microcontroller (ESP32): Dual-core microcontroller executing real-time threshold hysteresis, refractory blanking, transition timing, and JSON BLE notification framing.

2. Sensor Calibration & Force Conversion

2.1. 12-bit ADC Count to Vertical Force (kg)

The analog output of each FSR is sampled via a 12-bit voltage divider (0 to 4095 ADC counts). The calibrated transfer function maps raw sensor counts to vertical ground reaction force in kilograms:

Force_kg = max(0.0, (0.0125 * ADC_raw) - 13.75) Note: Sensor counts below 1100 ADC resolve to 0.0 kg, safely zeroing circuit noise, residual mechanical pre-load, and foot-off swing phase.

2.2. Superposition of Total Limb Reaction Force

At any instantaneous discrete time index t, the total vertical ground reaction force exerted by a limb is the sum of its regional heel and forefoot contact points:

F_total_L(t) = F_heel_L(t) + F_toe_L(t) F_total_R(t) = F_heel_R(t) + F_toe_R(t)

3. Step Detection & Event Filtering

3.1. Dual-Threshold Schmitt Trigger (Hysteresis)

To prevent contact chatter, jitter, and double-counting during stance load shifts, foot contact status is governed by a state machine with dual hysteresis thresholds:

• Contact Threshold (T_press): 1800 ADC counts (~8.75 kg equivalent force).

• Release Threshold (T_release): 1200 ADC counts (~1.25 kg equivalent force).

State(t) = Pressed, if ADC(t) > 1800 and State(t-1) == Released State(t) = Released, if ADC(t) <= 1200 and State(t-1) == Pressed State(t) = State(t-1), otherwise

3.2. Temporal Refractory Blanking Window

During dynamic gait, high-frequency impact spikes occur upon heel contact. A mandatory refractory blanking interval prevents multiple step triggers for a single stride:

Delta_t_foot = t_current - t_previous_step >= 300 ms

3.3. Bilateral Standing Transition Filter

When a user stands up or shifts balance while standing, both feet load almost simultaneously. If contralateral heel strike occurs within 150 ms of the ipsilateral strike:

if (abs(t_heel_L - t_heel_R) <= 150 ms): rollback_step_count() classify_as_static_double_stance() Prevents stationary posture adjustments from polluting dynamic ambulatory step counts.

4. Temporal Gait Parameters & Cadence

4.1. Cross-Limb Step Transitions (L -> R and R -> L)

Individual leg stride durations often remain identical even when an individual limps, because both extremities are locked to the same overall progression speed. True asymmetry manifests in the duration between contralateral strikes:

• Left-to-Right Transition (Delta_t_L2R): Elapsed time from Left foot initial contact (t_L) to Right foot initial contact (t_R).

• Right-to-Left Transition (Delta_t_R2L): Elapsed time from Right foot initial contact (t_R) to Left foot initial contact (t_L).

avg_L2R = (1 / N) * sum(Delta_t_L2R_k) for k=1..N (N=5) avg_R2L = (1 / N) * sum(Delta_t_R2L_k) for k=1..N (N=5)

4.2. Stride Duration (Gait Cycle Time)

Mean total stride duration (heel-to-heel duration for the same limb) is the sum of consecutive contralateral transitions:

Stride_Time_ms = avg_L2R + avg_R2L [or derived as (120,000 / Cadence)]

4.3. Cadence (Steps Per Minute)

Cadence is computed across standard 10-second evaluation intervals:

Cadence = clamp(Total_Steps_in_10s * 6, 0, 220) steps/min

5. Bilateral Symmetry & Clinical Limp Detection

5.1. Gait Symmetry Index (SI) and Symmetry Score

The Gait Symmetry Index measures the percentage discrepancy between left-to-right and right-to-left transition times:

mean_step = (avg_L2R + avg_R2L) / 2.0 Symmetry_Index (SI) = (abs(avg_L2R - avg_R2L) / mean_step) * 100.0% Symmetry_Score (%) = clamp(100.0 - (SI * 0.5), 0.0, 100.0) Normative clinical baseline: 92% to 100% in unimpaired adults. Lower values indicate asymmetric gait or antalgic limping.

5.2. Automated Limp Classification

A clinical limp diagnosis is derived from transition discrepancy delta = avg_L2R - avg_R2L:

if delta > +40 ms: classification = 'Left Leg Stance Dominant' elif delta < -40 ms: classification = 'Right Leg Stance Dominant' else: classification = 'Good Bilateral Symmetry'

6. Sub-Phasic Temporal Analysis (Phasogram)

Standard human walking splits into Stance Phase (~60% of gait cycle) and Swing Phase (~40% of gait cycle), accompanied by periods where both feet are grounded (Double Support, ~20%).

6.1. Continuous Stance Integration (50 Hz Telemetry)

For continuous session streams, cumulative active contact durations are integrated:

Stance_L (%) = clamp((Loaded_Time_L / Total_Session_Time) * 100.0, 45.0, 75.0) Stance_R (%) = clamp((Loaded_Time_R / Total_Session_Time) * 100.0, 45.0, 75.0) Double_Support (%) = clamp((Both_Loaded_Time / Total_Session_Time) * 100.0, 12.0, 40.0) Swing_L (%) = 100.0 - Stance_L (%) Swing_R (%) = 100.0 - Stance_R (%)

7. Dynamic Stability & Fall Risk Stratification

7.1. Gait Cycle Variability (Coefficient of Variation, CV%)

Cycle-to-cycle variability is a recognized diagnostic biomarker for dynamic balance control and ataxia:

CV_gait (%) = clamp((abs(avg_L2R - avg_R2L) / Stride_Time_ms) * 100.0 * 0.55, 1.1, 9.5) Healthy physiological baseline: CV < 2.5%. Values above 4.5% reflect balance instability.

7.2. Fall Risk & Mobility Evaluation Matrix

Fall Risk Level: 'LOW RISK': CV < 2.5% and Double_Support <= 25.0% 'MODERATE': CV <= 4.5% and Double_Support <= 30.0% 'HIGH RISK': Otherwise Mobility Status: 'FLUID & BALANCED': Symmetry >= 90.0% and CV < 3.0% 'FUNCTIONAL GAIT': Symmetry >= 80.0% 'COMPENSATION OBSERVED': Otherwise

8. Vertical Ground Reaction Force (vGRF) M-Wave Modeling

Human vertical ground reaction force follows a characteristic bimodal M-wave: an initial heel-strike shock peak, a midstance unweighting dip, and a terminal propulsive push-off peak.

8.1. Dual-Gaussian Physiological Model

For normalized stance progression u in [0.0, 1.0], force is modeled by dual Gaussians constrained by a sinusoidal boundary window W(u) = sin(pi * u):

W(u) = sin(pi * u) Heel_Peak(u) = F_target * 0.95 * exp(-((u - 0.22) / 0.13)^2) * (W(u) / 0.637) Toe_Peak(u) = F_target * 1.02 * exp(-((u - 0.72) / 0.14)^2) * (W(u) / 0.774) vGRF_total(u) = max(0.0, Heel_Peak(u) + Toe_Peak(u)) Heel strike peaks at 22% of stance; forefoot propulsion peaks at 72% of stance.

8.2. Bilateral Mirrored Rendering

InSoul charts render Left Foot upward (+y, Emerald Green) and Right Foot mirrored downward (-y, Amber-Orange), enabling immediate bilateral comparative visualization without trace collision.

8.3. 5-Point Discrete Gaussian Smoothing

y_i* = (0.06 * y_{i-2}) + (0.24 * y_{i-1}) + (0.40 * y_i) + (0.24 * y_{i+1}) + (0.06 * y_{i+2})

9. 6-Axis IMU Kinematics & Exercise Tracking

9.1. Accelerometer & Gyroscope Scaling

a_x = raw_Ax / 16384.0 g, a_y = raw_Ay / 16384.0 g, a_z = raw_Az / 16384.0 g w_x = raw_Gx / 131.0 deg/s, w_y = raw_Gy / 131.0 deg/s, w_z = raw_Gz / 131.0 deg/s

9.2. Foot Pitch Angle (Calf Raise / Incline Tilt)

The pitch inclination of the foot relative to the gravity vector:

theta_pitch = atan2(a_x, sqrt(a_y^2 + a_z^2)) * (180.0 / pi) normalized_tilt = clamp(theta_pitch / 45.0, -1.0, 1.0) smoothed_tilt(t) = smoothed_tilt(t-1) + 0.18 * (raw_tilt(t) - smoothed_tilt(t-1))

9.3. Toe-Standing / Plantarflexion Exercise State

ToeStandingActive = (F_toe_L > 1800 or F_toe_R > 1800) and (F_heel_L <= 1200 and F_heel_R <= 1200)

10. Summary Matrix of Gait Metrics & Reference Norms

| Gait Metric | Primary Formulation | Clinical Norm | Units |
| --- | --- | --- | --- |
| FSR to Force | max(0, 0.0125*ADC - 13.75) | Body Weight % | kg |
| Cadence | Total_Steps_10s * 6 | 90 - 120 | steps/min |
| Stride Time | avg_L2R + avg_R2L | 1000 - 1200 | ms |
| Symmetry Index | (¦L2R - R2L¦ / mean) * 100 | < 6.0% | % |
| Symmetry Score | clamp(100 - SI*0.5, 0, 100) | > 90.0% | % |
| Stance Phase % | (t_loaded / t_total) * 100 | 58% - 62% | % |
| Swing Phase % | 100% - Stance% | 38% - 42% | % |
| Double Support % | (t_both / t_total) * 100 | 18% - 24% | % |
| Gait Variability | (¦L2R - R2L¦ / Stride) * 55 | < 2.5% | % |
| A/P Roll Time | Heel to toe strike shift | 140 - 180 | ms |
| Pitch Tilt Angle | atan2(ax, sqrt(ay^2 + az^2)) | -45 to +45 | degrees |